HJJDS6031

Parameter description:	2

Appearance	3

Introduction to Interface Display	4

Key	5

Operation introduction	6

Press down CH:	6

Press down PARM:	7

Press down TRIG:	7

First press down HORI:	8

Second press down HORI:	8

First press down MENU:	9

Second press down MENU:	9

Third press down MENU:	10

Fourth press down MENU:	11

Fifth press down MENU:	11

Parameter description:

Channel:  					1

Bandwidth:					30MHz

real-time sampling rate     		200MSa/s

The storage depth         		2Kpts

Input coupling            		AC、DC

Input impedance          		1MΩ 25pF

The maximum input voltage 		40V (probe X1); 400V (probe X10) can be measured 220V

voltage; (probe X100) 2000V voltage can be measured

Probe attenuation             	1X、10X

Set the probe attenuation factor	1X、10X、100X

Sampling Method				Real-time sampling(10ns-50ms)/Scan sampling(100ms-5s)

Vertical Sensitivity 				10mV-5V (Probe 1X) 100mV-50V (probe 10X) (1,2.5,5 step)

Vertical accuracy				+/-3%

Vertical resolution				8bit

Horizontal scan range			10nS/div-5S/div(1,2.5,5 step)

Trigger Mode					Auto, Normal and Single

Trigger Type					Rising edge trigger, falling edge trigger

Automatic detection 			Support (50Hz-40MHz)

Cursor measurements			Support time and voltage cursors

Screen 						3.2-inch, 16-bit true color, TFT, 320 * 240

Battery						2800 + mA lithium battery (single cell about four Six of

continuous work)

Size 							195 * 99* 39(mm)

Appearance

Introduction to Interface Display

Key

Main buttons as follows

CH						Access Channel Control

PARM					Display parameters

Power					Turn on and off

AUTO					Automatic retrieval of waveforms

TRIG						Trigger control

HORI					Time-based correlation control

STOP/RUN				Stop start waveform display

MENU					Control menu

OK						Under the corresponding menu Function key

↑						Under the corresponding menu Function key

↓						Under the corresponding menu Function key

←						Under the corresponding menu Function key

→						Under the corresponding menu Function key

# F1						Under the corresponding menu Function key

F2						Under the corresponding menu Function key

F3						Under the corresponding menu Function key

Operation introduction

## Press down CH:

Function key description

↑						Voltage shift control

↓						Voltage shift control

←						Waveform Horizontal Position Control

→						Waveform Horizontal Position Control

F1						Open Closed Channel Display

F2						probe x1,x10,x100

F3						AC and DC

## Press down PARM:

Function key description

↑						Voltage shift control

↓						Voltage shift control

←						Waveform Horizontal Position Control

→						Waveform Horizontal Position Control

F1						Open Closed Channel Display

F2						probe x1,x10,x100

F3						AC and DC

## Press down TRIG:

Function key description

↑						Trigger position shift

↓						Trigger position shift

←						Trigger position shift

→						Trigger position shift

F1						Rising edge and falling edge trigger

F2						invalid

F3						Select trigger type automatically normal single

## First press down HORI:

Function key description

↑						Changing time base

↓						Changing time base

←						Change the position of the time base arrow

→						Change the position of the time base arrow

F1						Open and close measuring scale

F2						The object of scale measurement

F3						Type Voltage Time Measured by Scale

## Second press down HORI:

Function key description

↑						Position of scale 2 of the table

↓						Position of scale 2 of the table

←						Position of scale 1 of the table

→						Position of scale 1 of the table

F1						invalid

F2						invalid

F3						invalid

## First press down MENU:

Function key description

↑						invalid

↓						invalid

←						invalid

→						invalid

F1						Label of stored waveform A or B

F2						Source of stored waveform

F3						Whether to display stored waveforms

## Second press down MENU:

Function key description

↑						invalid

↓						invalid

←						invalid

→						invalid

F1						Change the brightness level of backlight

F2						Language Selection of Interfaces

F3						Auto-calibration

## Third press down MENU:

Function key description

↑						invalid

↓						invalid

←						invalid

→						invalid

F1						Turn on and off button sound

F2						Turn on and off automatic shutdown

F3						Turn on and off automatic power saving

## Fourth press down MENU:

Function key description

↑						invalid

↓						invalid

←						invalid

→						invalid

F1						invalid

F2						Setting up screenshots

F3						Factory Default

## Fifth press down MENU:

Function key description

↑						invalid

↓						invalid

←						invalid

→						invalid

F1						invalid

F2						invalid

F3						Turn on the inversion of the closed waveform

OK Explain:

Used for screenshots when screenshots are turned on

Used to store waveforms when screenshots are turned off